Hellenic Accreditation System

Annex F2/9 to the Certificate No. 887-4

SCOPE of ACCREDITATION

of the

Testing & Calibration Laboratory

of

METRISIS CALIBRATIONS P.C.

for the performance of calibrations

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass measurements |
| Standard weights |  |  | Applied guideline: OIML R111:2004 Capability to calibrate standard weights up to OIML class: |
|  | 1 mg to 2 mg | 8,7 μg | F2 |
|  | 5 mg to 10 mg | 8,9 μg | F2 |
|  | 20 mg to 50 mg | 8,9 μg | F1 |
|  | 100 mg to 200 mg | 9,0 μg | F1 |
|  | 500 mg | 9,1 μg | F1 |
|  | 1 g | 9,4 μg | E2 |
|  | 2 g | 9,7 μg | E2 |
|  | 5 g | 10 μg | Ε2 |
|  | 10 g | 11 μg | Ε2 |
|  | 20 g | 12 μg | Ε2 |
|  | 50 g | 15 μg | Ε2 |
|  | 100 g | 87 μg | F1 |
|  | 200 g | 92 μg | Ε2 |
|  | 500 g | 0,92 mg | F2 |
|  | 1 kg | 1,0 mg | F1 |
|  | 2 kg | 1,9 mg | F1 |
|  | 5 kg | 9,0 mg | F2 |
|  | 10 kg | 83 mg | M1 |
|  | 20 kg | 83 mg | F2 |
| Non-automatic weighing instruments | 1 mg to 20 mg | 0,004 mg | Applied guideline: EURAMET/cg-18/v.04:2015 Standard weights used, OIML class: E1: 1 g to 611 g E2: 1 mg to 38,5 kg F1: 1 mg to 50,8 kg F2: 1 mg to 9,1 kg M1: 5 kg to 360 kg M2: 50 g to 340 kg M3: 50 g to 460 kg Calibration is performed on-site. |
|  | 50 mg | 0,005 mg |  |
|  | 100 mg | 0,007 mg |  |
|  | 200 mg | 0,008 mg |  |
|  | 500 mg | 0,011 mg |  |
|  | 1 g | 0,013 mg |  |
|  | 2 g | 0,016 mg |  |
|  | 5 g | 0,021 mg |  |
|  | 10 g | 0,027 mg |  |
|  | 20 g | 0,033 mg |  |
|  | 50 g | 0,040 mg |  |
|  | 100 g | 0,067 mg |  |
|  | 200 g | 0,13 mg |  |
|  | 500 g | 0,33 mg |  |
|  | 1 kg | 2,1 mg |  |
|  | 2 kg | 4,0 mg |  |
|  | 5 kg | 11 mg |  |
|  | 10 kg | 21 mg |  |
|  | 20 kg | 40 mg |  |
|  | 50 kg | 0,33 g |  |
|  | 100 kg | 3,7 g |  |
|  | 200 kg to360 kg | 6, 7 x 10-5 |  |
|  | 500 kg | 52 g |  |
|  | 700 kg | 92 g |  |
|  | 1000 kg | 0,29 kg |  |
|  | 1160 kg | 0,40 kg |  |
|  | 1160 kg to10440 kg | (3,44 x 10-4 x m + 1,2 x N) kg m: load of standard weights in kg Ν: number of substitution loads | Applied guideline: EURAMET/cg-18/v.04:2015 using substitution loads of known material and density supplied by the customer according to the recommendations of the laboratory. Standard weights used, OIML class: M1: 360 kg M2: 340 kg M3: 460 kg Calibration is performed on-site. |
| Vehicle suspension testers | 0 kg to 1160 kg per wheel | 0,090 % | Calibration according to manufacturer guidelines: CARTEC, BOSCH, ΑΤΤ, MAHA, BEISSBARTH και CORGHI Calibration is performed on-site. |
| Exhaust Gas analyzers calibration |
| Exhaust gas analyzers class 0, I and II for gasoline operated vehicles, portal exhaust gas analyzers | C3H8 : 2000 x 10-6 vol 600 x 10-6 vol 200 x 10-6 vol | 0,60 % 0,74 % 0,95 % | Applied standards: OIML R 99-1&2:2008 Calibration can also be performed on-site. |
|  | CO : 3,5 % vol 1,0 % vol 0,5 % vol | 0,34 % 0,29 % 0,34 % |  |
|  | CO2 : 14 % vol 10 % vol 6 % vol | 0,46 % 0,46 % 0,47 % |  |
|  | O2 : 0,5 % vol to 20 % vol | 1,0 % |  |
| Exhaust gas analyzers for diesel operated vehicles | Absorption: 0% to 100% | 0,80 % | Calibration according to directive 72/306/EEC, Annex VII, par. 3.6. Available standard reference filters: HELMA, MAHA, SENSORS/ CARTEC, TEXA, BOSCH, TECHNOTEST Calibration can also be performed on-site. |
|  | Absorption coefficient (opacity): 0 m-1 to 9,99 m-1 | 0,040 m-1 |  |
| Dimensional measurements |
| Vehicle side slip testers | -20 mm to +20 mm | 0,010 mm | Calibration according to manufacturer guidelines: CARTEC, BOSCH, ΑΤΤ, MAHA, BEISSBARTH, CORGHI Calibration is performed on-site. |
| Force measurements |
| Vehicle brake force testers |  |  | Calibration according to manufacturer guidelines: |
|  | 0 kN to 40 kΝ | 0,90 % | CARTEC |
|  | 0 kN to 5 kΝ 5 kN to 40 kN | 0,40 % 0,20 % | ΜΑΗΑ |
|  | 0 kN to 40 kN | 0,40 % | BOSCH/ΑΤΤ |
|  | 0 kN to 40 kΝ | 0,40 % | CORGHI |
|  | 0 kN to 40 kN | 0,30 % | BEISSBARTH |
|  |  |  | Calibration is performed on-site. |
| Photometric measurements |
| Vehicle headlight testers | Illuminance: 0 to 240 lux | 2,6 % | Calibration according to manufacturer guidelines: CARTEC, BOSCH, MAHA, BEISSBARTH, TECNOLUX & TECNOIL, TECNOCOLOR, CORGHI Calibration is performed on-site. |
|  | Inclination of incident light beam: 0% 2% (vertically) 2% (horizontally) | 0,032 % 0,056 % 0,040 % |  |
| Temperature measurements |
| Temperature / temperature controlled climatic chambers of volume up to 2000 l (ovens, incubators, autoclaves, refrigerators, freezers), isothermal blocks, PCR | -80 °C to -20 °C | 0,16 °C | Applied guideline: DKD-R 5-7:2018 For temperature ranges -80 °C up to -20 °C and 180 °C up to 400 °C calibration is performed at different positions of the chamber according to method (C) of the guideline. Calibration can also be performed on-site. |
|  | -20 °C to 0 °C | 0,15 °C |  |
|  | 0 °C to 150 °C | 0,14 °C |  |
|  | 150 °C to 180 °C | 0,22 °C |  |
|  | 180 °C to 250 °C | 0,25 °C |  |
|  | 250 °C to 400 °C | 1,1 °C |  |
| Direct reading thermometers, temperature recorders with external sensor | -20 °C to 0 °C | 0,05 °C | Applied guideline: DKD-R 5-1:2018 Calibration can also be performed on-site. |
|  | 0 °C to 60 °C | 0,04 °C |  |
|  | 60 °C to 120 °C | 0,03 °C |  |
|  | 120 °C to 150 °C | 0,05 °C |  |
|  | 150 °C to 180 °C | 0,17 °C |  |
| Temperature recorders | -10 °C to 0 °C | 0,20 °C | Applied guideline: DKD-R 5-1:2018 |
|  | 0 °C to 45 °C | 0,15 °C |  |
|  | 45 °C to 70 °C | 0,23 °C |  |
| Liquid in-glass thermometers of partial/total immersion | -20 °C to 0 °C | 0,05 °C | Applied guidelines: COFRAC LAB GTA 08:2018 ISO 386:1977 ASTM E77:2014 Calibration in isothermal bath. Calibration can also be performed on-site. |
|  | 0 °C to +60 °C | 0,04 °C |  |
|  | 60 °C to 120 °C | 0,03 °C |  |
|  | 120 °C to 150 °C | 0,05 °C |  |
|  | 150 °C to 180 °C | 0,17 °C |  |
| Volume measurements |
| Piston operated pipettes | 10 μl | 0,35% | Applied guideline: EURAMET/cg-19/ v.3.0 2018 |
|  | 20 μl | 0,20% |  |
|  | 50 μl | 0,13% |  |
|  | 100 μl έως 10 ml | 0.12% |  |
| Piston volumetric apparatus (dispensers, digital burettes, titrators) | 1ml to 100 ml | 0,13 % |  |
| Bulb pipettes, volumetric flasks, cylinders | 1 ml | 0,20 % |  |
|  | 2 ml to 5 ml | 0,10 % |  |
|  | 10 ml to 15 ml | 0,050 % |  |
|  | 20 ml to 25 ml | 0,040 % |  |
|  | 50 ml to 100 ml | 0,030 % |  |
|  | 150 ml to 500 ml | 0,020 % |  |
|  | 1000 ml to 2000 ml | 0,030 % |  |
| Spectrophotometric measurements |
| Absorption / Spectrophotometers of visible and UV radiation, in the range: 235 nm to 695 nm | 0 to 0,65 AU (UV) | 5,2 mAU** | Applied standard: ASTM E 275-08(2022) Calibration can also be performed on-site. |
|  | 0 AU to 0,25 AU (visible) | 2,8 mAU** |  |
|  | 0,25 AU to 0,50 AU (visible) | 3,2 mAU** |  |
|  | 0,50 AU to 1,00 AU (visible) | 3,9 mAU** |  |

*   Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not  accompanied by a unit, it is relative.

** Uncertainty values are estimated assuming a spectrophotometer having a resolution of 0,1mAU.

Site of assessment: Permanent premises, 95 Lambrou Katsoni str, 183 44, Moschato, Piraeus.

Approved Signatories: C. Konstantopoulos, G. Fardis.

This Scope of Accreditation replaces the previous one dated 16.03.2022.

The Accreditation Certificate No. 887-4, to ELOT EN ISO/IEC 17025:2017, is valid until 29.09.2025.

Athens, 16.05.2022

Christos Nestoras

CEO of ESYD