Hellenic Accreditation System

Annex F2/18 to the Certificate No. 90-7

SCOPE of ACCREDITATION of the Calibration Laboratory of the Innovation Hub (DKK) of PPC S.A.

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| DC Voltage (Measurement)/ Voltage Standards, Voltage Calibrators and Sources |  | parts per 106 output + μV | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [1 μV to 100 μV] | 0 + 0.029 |  |
|  | (100 μV to 1 mV] | 0 + 0.042 |  |
|  | (1 mV to 10 mV] | 0 + 0.25 |  |
|  | (10 mV to 120 mV] | 5.0 + 0.3 |  |
|  | (120 mV to 1.2 V] | 4.0 + 0.8 |  |
|  | 1.018 V | 0 + 0.74 |  |
|  | (1.2 V to 12 V] | 4.0 + 0.5 |  |
|  | 10 V | 0 + 0.38 |  |
|  | (12 V to 120 V] | 6.0 + 30 |  |
|  | (120 V to 1050 V] | 6.0 + 100 |  |
|  | [1 kV to 10 kV] | 136.0 + 0 |  |
| DC Voltage (Generation)/ Voltage Measuring Equipment | [1 μV to 100 μV] | 0 + 0.01 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | (100 μV to 1 mV] | 0 + 0.04 |  |
|  | (1 mV to 5 mV] | 0 + 0.17 |  |
|  | (5 mV to 10 mV] | 0 + 0.34 |  |
|  | (10 mV to 220 mV] | 6.5 + 0.8 |  |
|  | (220 mV to 2.2 V] | 4.0 + 0.8 |  |
|  | 1.018 V | 0 + 0.74 |  |
|  | (2.2 V to 12 V] | 3.0 + 2.5 |  |
|  | 10 V | 0 + 0.38 |  |
|  | (12 V to 22 V] | 3.0 + 4 |  |
|  | (22 V to 220 V] | 4.0 + 40 |  |
|  | (220 V to 1100 V] | 6.0 + 400 |  |

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| AC Voltage (Measurement)/ Voltage Standards, Voltage Calibrators and Sources | (12 mV to 12 V] [1 Hz to 40 Hz] (40 Hz to 1 kHz] (1 kHz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 1 MHz] (1 MHz to 2 MHz] (2 MHz to 4 MHz] (4 MHz to 8 MHz] (8 MHz to 10 MHz] | 0.007 + 0.004 0.007 + 0.002 0.014 + 0.002 0.03 + 0.002 0.08 + 0.002 0.3 + 0.01 1 + 0.01 1.5 + 0.01 4 + 0.07 4 + 0.08 15 + 0.1 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | (12 V to 120 V] [1 Hz to 40 Hz] (40 Hz to 1 kHz] (1 kHz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 1 MHz] | 0.02 + 0.004 0.02 + 0.002 0.02 + 0.002 0.035 + 0.002 0.12 + 0.002 0.4 + 0.01 1.5 + 0.01 |  |
|  | (120 V to 700 V] [1 Hz to 40 Hz] (40 Hz to 1 kHz] (1 kHz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] [0.7 kV to 3 kV) [50 Hz to 60Hz] [3 kV to 7 kV] [50 Hz to 60Hz] | 0.04 + 0.004 0.04 + 0.002 0.06 + 0.002 0.12 + 0.002 0.3 + 0.002 % Reading + V 0.577 + 0 0.211 + 0 |  |

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| AC Voltage (Generation)/ Voltage Measuring Equipment |  | parts per 106 output + μV | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [220 μV to 2.2 mV) [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] | 230 + 4 87 + 4 77 + 4 190 + 4 480 + 5 1000 + 10 1300 + 20 2600 + 20 |  |
|  | [2.2 mV to 22 mV) [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] | 230 + 4 87 + 4 77 + 4 190 + 4 480 + 5 1000 + 10 1300 + 20 2600 + 20 |  |
|  | [22 mV to 220 mV] [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] | 230 + 12 87 + 7 77+ 7 190 + 7 440 + 17 800 + 20 1300 + 25 2600 + 45 |  |
|  | (220 mV to 2.2 V] [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] | 230 + 40 85 + 15 42 + 8 73 + 10 107 + 30 380 + 80 950 + 200 1600 + 300 |  |

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| (continued) AC Voltage (Generation) / Voltage Measuring Equipment |  | parts per 106 output + μV | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | (2.2 V to 22 V] [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] (22 V to 220 V] [10 Hz to 20 Hz] (20 Hz to 40 Hz] (40 Hz to 20 kHz] (20 kHz to 50 kHz] (50 kHz to 100 kHz] (100 kHz to 300 kHz] (300 kHz to 500 kHz] (500 kHz to 1 MHz] (220 V to 1100 V] [50 Hz to 1 kHz] (1 kHz to 20 kHz] (20 kHz to 30 kHz] to 750 V] (30 kHz to 50 kHz] (50 kHz to 100 kHz] | 230 + 400 85 + 150 42 + 50 73 + 100 97 + 200 270 + 600 900 + 2000 1400 + 3200 230 + 4000 85 + 1500 50 + 600 77 + 1000 140 + 2500 850 + 16000 4300 + 40000 7500 + 80000 65 + 3500 135 + 6000 440 + 11000 440 + 11000 1600 + 45000 |  |
| DC Current (Measurement) / Current Calibrators and Sources |  | parts per 106 Reading + parts per 106 Range | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [10 nA to 120 nA) | 30 + 400 |  |
|  | [0.12 μA to 1.2 μA) | 20 + 40 |  |
|  | [1.2 μA to 12 μA) | 20 + 10 |  |
|  | [12 μA to 120 μA) | 20 + 8 |  |
|  | [0.12 mΑ to 1.2 mΑ) | 20 + 5 |  |
|  | [1.2 mΑ to 12 mΑ) | 20 + 5 |  |
|  | [12 mΑ to 120 mΑ) | 35 + 5 |  |
|  | [0.12 Α to 1.05 Α) [1.05 A to 200 A] | 110 + 10 240 + 0 |  |
| DC Current (Generation)/ Current Measuring Equipment |  | parts per 106 output + nA | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [10 μA to 220 μA) | 37 + 6 |  |
|  | [0.22 mΑ to 2.2 mΑ) | 33 + 7 |  |
|  | [2.2 mΑ to 22 mΑ) | 33 + 40 |  |
|  | [22 mΑ to 220 mΑ) | 42 + 700 |  |
|  | [0.22 Α to 2.2 Α) | 70 + 12000 |  |
|  | [2.2 Α to 11 Α] | 350 + 480000 |  |
| Current Clamp Type Measuring Equipment |  | % Reading + mA |  |
|  | [10 Α to 1000 Α] | 0.326 + 0 |  |

| AC Current (Measurement)/ Current Calibrators and Sources | [10 μA to 120 μA) [10 Hz to 20 Hz) [20 Hz to 45 Hz) [45 Hz to 1 kHz] | % Reading + % Range 0.4 + 0.03 0.15 + 0.03 0.06 + 0.03 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
| --- | --- | --- | --- |
|  | [0.12 mA to 120 mA) [10 Hz to 20 Hz) [20 Hz to 45 Hz) [45 Hz to 100 Hz) [100 Hz to 5 kHz) [5 kHz to 20 kHz) [20 kHz to 50 kHz) [50 kHz to 100 kHz] | 0.4 + 0.02 0.15 + 0.02 0.06 + 0.02 0.03 + 0.02 0.06 + 0.02 0.4 + 0.04 0.55 + 0.15 |  |
|  | [120 mA to 1.05 Α] [10 Hz to 20 Hz) [20 Hz to 45 Hz) [45 Hz to 100 Hz) [100 Hz to 5 kHz) [5 kHz to 20 kHz) [20 kHz to 50 kHz] [20 A to 100 A) 50 Hz [100 A to 700 A] 50 Hz (700 A to 1000 A] 50 Hz [1 kA to 10 kA] [50 to 60 Hz] | 0.4 + 0.02 0.16 + 0.02 0.08 + 0.02 0.1 + 0.02 0.3 + 0.02 1 + 0.04 3.5 + 0 0.76 + 0 1.89 + 0 1.10 + 0 |  |
| AC Current (Generation)/ Current Measuring Equipment |  | parts per 106 output + nA | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [0.22 mA to 2.2 mA) [10 Hz to 20 Hz) [20 Hz to 40 Hz) [40 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] | 240 + 16 150 + 10 115 + 8 270 + 12 1000 + 65 |  |
|  | [2.2 mA to 22 mA) [10 Hz to 20 Hz) [20 Hz to 40 Hz) [40 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] | 240 + 40 150 + 35 115 + 35 190 + 110 1000 + 650 |  |
|  | [0.22 mA to 2.2 mA) [10 Hz to 20 Hz) [20 Hz to 40 Hz) [40 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] | 240 + 400 150 + 350 115 + 350 190 + 550 1000 + 5000 |  |

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| (continued) AC Current (Generation)/ Current Measuring Equipment |  | parts per 106 output + μA | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [22 mA to 220 mA) [10 Hz to 20 Hz) [20 Hz to 40 Hz) [40 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] | 240 + 4 150 + 3.5 115 + 2.5 190 + 3.5 1000 + 10 |  |
|  | [220 mA to 2.2 A) [20 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] | 250 + 35 420 + 80 6500 + 160 |  |
|  | [2.2 A to 11 A] [40 Hz to 1 kHz) [1 kHz to 5 kHz) [5 kHz to 10 kHz] [10 A to 100 A) 50 Hz (50 Hz to 400 Hz] [100 A to 400 A] 50 Hz (50 Hz to 400 Hz] (400 A to 1000 A] 50 Hz | 440 + 170 900 + 380 3500 + 750 % Reading + A 3.2 + 0 8.4 + 0 0.32 + 0 0.86 + 0 0.083 + 0 |  |
| Current Clamp Type Measuring Equipment |  |  |  |
| Resistance DC (Generation)/ Resistance Measuring Equipment |  | parts per 106 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | 0 Ω | 40 μΩ |  |
|  | 10 μΩ | 42 |  |
|  | 100 μΩ | 23 |  |
|  | 1 mΩ | 19 |  |
|  | 10 mΩ | 13 |  |
|  | 100 mΩ | 13 |  |
|  | 1 Ω | 85 |  |
|  | 1.9 Ω | 85 |  |
|  | 10 Ω | 22 |  |
|  | 19 Ω | 22 |  |
|  | 100 Ω | 9.5 |  |
|  | 190 Ω | 9.5 |  |
|  | 1 kΩ | 8.0 |  |
|  | 1.9 kΩ | 8.0 |  |
|  | 10 kΩ | 85 |  |
|  | 19 kΩ | 8.0 |  |
|  | 100 kΩ | 10 |  |
|  | 190 kΩ | 10 |  |
|  | 1 ΜΩ | 17 |  |
|  | 1.9 ΜΩ | 18 |  |
|  | 10 ΜΩ | 34 |  |
|  | 19 ΜΩ | 42 |  |
|  | 100 ΜΩ | 100 |  |

| Parameters/ Calibration Item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Resistance DC (Measurement)/ Standard Resistors |  | parts per 106 |  |
|  | [1 μΩ to 10 μΩ) | 3000 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [10 μΩ to 100 μΩ) | 380 |  |
|  | [100 μΩ to 1 mΩ) | 180 |  |
|  | [1 mΩ to 10 mΩ) | 60 |  |
|  | [10 mΩ to 100 mΩ) | 10 |  |
|  | [100 mΩ to 1 Ω] | 7.3 |  |
|  | (1 Ω to 10 Ω] | 7.3 |  |
|  | (10 Ω to 100 Ω) | 11 |  |
|  | [100 Ω to 1 kΩ) | 11 |  |
|  | [1 kΩ to 10 kΩ] | 11 |  |
|  | (10 kΩ to 100 kΩ] | 11 |  |
|  | (100 kΩ to 1 ΜΩ] | 10 |  |
|  | (1 ΜΩ to 10 ΜΩ] | 38 |  |
|  | (10 ΜΩ to 100 ΜΩ] | 79 |  |
|  | (100 ΜΩ to 1 GΩ] | 5900 |  |
| Resistance AC (Measurement)/ Standard Resistors | 10 Ω | [220 μΑ to 2.2 mA] | [50 Hz to 125 Ηz] | 290 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | [100 Ω to 10 kΩ] |  |  | 180 |  |
|  | 1 Ω | [2.2 mA to 22 mA] |  | 290 |  |
|  | [10 Ω to 1 kΩ] |  |  | 180 |  |
|  | 100 mΩ | [22 mA to 220 mA] |  | 240 |  |
|  | [1 Ω to 100 Ω] |  |  | 180 |  |
|  | 10 mΩ | [220 mA to 2.2 A] |  | 340 |  |
|  | [100 mΩ to 10 Ω] |  |  | 300 |  |
|  | [10 mΩ to 100 mΩ] | [2.2 A to 11 A] |  | 470 |  |
| Resistance AC (Generation)/ Resistance Measuring Equipment | 10 mΩ | 10 A | [50 Hz to 125 Ηz] | 1600 | Calibration according to (i) EURAMET cg-15, (ii) standards of lab, (iii) lab internal procedures |
|  | 100 mΩ | 3 A |  | 840 |  |
|  | 1 Ω | 1.4 A |  | 12000 |  |
|  | 10 Ω | 100 mA |  | 1300 |  |
|  | 100 Ω | 20 mA |  | 680 |  |
|  | 1 kΩ | 10 mA |  | 680 |  |
|  | 10 kΩ | 3 mA |  | 790 |  |

* Where the expanded uncertainty (with 95 % coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration Measurement Capability (CMC) includes the measured quantity, the measurement range and the measurement uncertainty, expressing the minimum measurement uncertainty which can be achieved in a calibration.

Permanent laboratory premises address: 9 Leondariou Str., 15351 Kantza, Pallini, Attiki, Greece.

Approved signatories: M. Bomboulos, D. Kaimaras, M. Valsamakis, E. Thirios, A. Petrakos.

This Scope of Accreditation replaces the previous one dated 14.03.2023.

The Accreditation Certificate No. 90-7, to ELOT EN ISO/IEC 17025:2017, is valid until 02.07.2027.

Athens, 07.09.2023

Christos Nestoras

CEO of ESYD