Hellenic Accreditation System

Annex F2/3 to the Certificate No. 1035-2

SCOPE of ACCREDITATION

of the

Calibration Laboratory

MATELEC METROLOGY DIVISION

| Measurand / Calibration item | Range of measurement | Uncertainty * (Confidence level 95%) | Remarks |
| --- | --- | --- | --- |
| Electrical measurements |
| DC Voltage / DC voltage meters | (0 – 0.2] V | 4 μV + 17·10-6·U | Calibration using standard multifunction calibrator according to guide EURAMET cg-15 v.3.0 (2015). |
|  | (0.2 - 2] V | 8μV + 10.4·10-6·U |  |
|  | (2 – 20.2] V | 29 μV + 9.2·10-6·U |  |
|  | (20.2 - 202] V | 280 μV + 14·10-6·U |  |
|  | (202 - 1020] V | 2.8 mV + 14·10-6·U |  |
|  |  | U: measured voltage |  |
| DC Current / DC current meters | (0 – 0.2] mA | 0.02 μA + 2·10-4·I | Calibration using standard multifunction calibrator according to guide EURAMET cg-15 v.3.0 (2015). |
|  | (0.2 – 2] mA | 0.04 μA + 6·10-4·I |  |
|  | (2 – 20.2] mA | 0.3 μA + 0.6·10-4·I |  |
|  | (20.2 – 202] mA | 3 μA + 0,6·10-4·I |  |
|  | (0.2 – 2] A | 35 μA + 1,6·10-4·I |  |
|  | (2 – 20.2] A | 400 μA + 4·10-4·I |  |
|  | (20.2 – 30] A | 520 μA + 6·10-4·I |  |
|  |  | I: measured current |  |
| Resistance/Resistance measuring instruments (4 / 2 wire) | (0 – 1] Ω | 6 mΩ + 2·10-4·R | Calibration using standard multifunction calibrator according to guide EURAMET cg-15 v.3.0 (2015). |
|  | (1 – 10] Ω | 6 mΩ + 2·10-4·R |  |
|  | (10 – 100] Ω | 6 mΩ + 0.6·10-4·R |  |
|  | (0.1 – 1] kΩ | 50 mΩ + 0.5·10-4·R |  |
|  | (0.1 - 10] kΩ | 0.5 Ω + 0.5·10-4·R |  |
|  | (10 - 100] kΩ | 5 Ω + 0.5·10-4·R |  |
|  | (0.1 - 1] MΩ | 46 Ω + 2·10-4·R |  |
|  | (1 - 10] MΩ | 500 Ω + 4·10-4·R |  |
|  | (10 - 100] MΩ | 5 kΩ + 6·10-3·R |  |
|  | (0.1 - 1] GΩ | 50 kΩ + 20·10-3·R |  |
|  |  | R: measured resistance |  |
| AC Voltage / AC voltage measuring instruments | (0 - 0.2] V / [10 - 44] Hz | 19 μV + 10·10-4·U | Calibration using standard multifunction calibrator according to guide EURAMET cg-15 v.3.0 (2015). |
|  | / [45 - 999] Hz | 19 μV + 2·10-4·U |  |
|  | / [1 - 20) kHz | 39 μV + 2·10-4·U |  |
|  | / [20 - 100) kHz | 46 μV + 20·10-4·U |  |
|  | (0.2 - 2] V / [10 - 44] Hz | 220 μV + 6·10-4·U |  |
|  | / [45 - 999] Hz | 150 μV + 2·10-4·U |  |
|  | / [1 - 20) kHz | 220 μV + 2.4·10-4·U |  |
|  | / [20 - 100) kHz | 350 μV + 7.6·10-4·U |  |
|  | (2 V - 20] V / [10 - 44] Hz | 1.9 mV + 8·10-4·U |  |
|  | / [45 - 999] Hz | 1.3 mV + 2·10-4·U |  |
|  | / [1 - 20) kHz | 1.9 mV + 2.4·10-4·U |  |
|  | / [20 - 100) kHz | 3.5 mV + 8·10-4·U |  |
|  | (20 V - 202] V / [30 – 44] Hz | 25 mV + 6·10-4·U |  |
|  | / [45 – 999] Hz | 15 mV + 1.8·10-4·U |  |
|  | / [1 - 10) kHz | 19 mV + 2·10-4·U |  |
|  | / [10 – 40] kHz | 35 mV + 4·10-4·U |  |
|  | (202 - 1020] V / [30 – 44] Hz | 240 mV + 6.4·10-4·U |  |
|  | / [45 – 999] Hz | 910 mV + 3·10-4·U |  |
|  |  | U: measured voltage |  |
| AC Current / AC current measuring instruments | [20 - 202] μA / [10 – 44] Hz | 0.28 μA + 0.3·10-2·I | Calibration using standard multifunction calibrator according to guide EURAMET cg-15 v.3.0 (2015). |
|  | / [45 – 999] Hz | 0.18 μA + 0.12·10-2·I |  |
|  | / [1 – 10] kHz | 0.28 μA + 1.0·10-2·I |  |
|  | (0.2 - 2] mA / [10 – 44] Hz | 0.29 μA + 0.3·10-2·I |  |
|  | / [45 – 999] Hz | 0.3 μA + 0.10·10-2·I |  |
|  | / [1 – 10] kHz | 0.4 μA + 0.6·10-2·I |  |
|  | (2 -20.2] mA / [10 – 44] Hz | 4 μA + 0.2·10-2·I |  |
|  | / [45 – 999] Hz | 3 μA + 0.10·10-2·I |  |
|  | / [1 – 10] kHz | 4 μA + 0.29·10-2·I |  |
|  | (20.2 - 202] mA / [10 – 44] Hz | 30 μA + 0.3·10-2·I |  |
|  | / [45 – 999] Hz | 30 μA + 0.1·10-2·I |  |
|  | / [1 – 10] kHz | 50 μA + 0.6·10-2·I |  |
|  | (0.2 - 2] A / [10 – 44] Hz | 350 μA + 0.3·10-2·I |  |
|  | / [45 – 999] Hz | 250 μA + 0.1·10-2·I |  |
|  | / [1 – 10] kHz | 460 μA + 0.6·10-2·I |  |
|  | (2 -30] A / [30- 44] Hz | 3.5 mA + 0.24·10-2·I |  |
|  | / [45 – 99] Hz | 2.4 mA + 0.07·10-2·I |  |
|  |  | I: measured current |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative

Site of assessment: Permanent laboratory premises, Ghorfine-Jbeil-Lebanon, PO-BOX 10 Jbeil

Approved Signatories: T. A. Fadel

This Scope of Accreditation replaces the previous one dated September 9th, 2020.

The Accreditation Certificate No. 1035, to ELOT EN ISO/IEC 17025:2017, is valid until 19.04.2024.

Athens, October 14th, 2022

Christos Nestoras

CEO of ESYD