Hellenic Accreditation System

Annex F2/19 to the Certificate No. 426-6 SCOPE of ACCREDITATION of the Laboratory of TOTAL Q CALIBRATION LABORATORIES S.A. for the performance of calibrations

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Volume measurements |
| Water volume / Piston pipettes | [0,1 to 1) μL | 5,0 10 % | Internal procedure tQ-VOLUME-01 Gravimetric method based on ISO 8655-6 (2022) ISO/TR 20461 (2023) |
|  | [1 to 2] μL | 2,0 % | Gravimetric method according to: ISO 8655-6 (2022) ISO/TR 20461 (2023) |
|  | (2 to 50] μL | 0,50 % |  |
|  | (50 to 500] μL | 0,19 % |  |
|  | (0,5 to 2,0] mL | 0,15 % |  |
|  | (2,0 to 10,0] mL | 0,14 % |  |
| Water volume / Piston volumetric apparatus (dispensers, digital burettes, titrators | 0,1 mL to 100 mL | 0,05 % | Internal procedure tQ-VOLUME-01 Gravimetric method based on ISO 8655-6 (2022) ISO/TR 20461 (2023) |
| Water volume / Volumetric flasks | [5 to 100] mL | 0,05 % | Gravimetric method according to ISO 8655-6 (2022) ISO 4787 (2021) ISO/TR 20461 (2023) |
|  | (100 to 2000] mL | 0,012 % |  |
| Water volume / Bulb pipettes | [1 to 10] mL | 0,05 % | Gravimetric method according to ISO 8655-6 (2022) ISO 4787 (2021) ISO/TR 20461 (2023) |
|  | (10 to 100] mL | 0,02 % |  |
| Water volume / Other volumetric apparatus of «In» and «Ex» type (one-mark and graduated volumetric glassware, plastic and metallic proving tanks) | [0,1 to 1) μL | 5,0 % | Internal procedure tQ-VOLUME-01 Gravimetric method based on ISO 8655-6 (2022) ISO/TR 20461 (2023) |
|  | [1 to 2] μL | 2,0 % |  |
|  | (2 to 50] μL | 0,50 % |  |
|  | (50 to 500] μL | 0,15 % |  |
|  | (0,5 to 2,0] mL | 0,10 % |  |
|  | (2 to 10,0] mL | 0,07 % |  |
|  | (10,0 to 2500] mL | 0,05 % |  |
| Mass measurements |
| Mass / Standard weights |  |  | Standard weights up to OIML R111-1: 2004 class: |
|  | [1 to 5] mg | 2,0 μg | Ε2 |
|  | 10 mg | 2,5 μg | Ε2 |
|  | 20 mg | 3,3 μg | Ε2 |
|  | 50 mg | 4,0 μg | Ε2 |
|  | 100 mg | 5,0 μg | Ε2 |
|  | 200 mg | 6,5 μg | Ε2 |
|  | 500 mg | 8,0 μg | Ε2 |
|  | 1 g | 10 μg | Ε2 |
|  | 2 g | 13 μg | Ε2 |
|  | 5 g | 16 μg | Ε2 |
|  | 10 g | 20 μg | Ε2 |
|  | 20 g | 26 μg | Ε2 |
|  | 50 g | 33 μg | Ε2 |
|  | 100 g | 53 μg | Ε2 |
|  | 200 g | 0,10 mg | Ε2 |
|  | 500 g | 0,25 mg | Ε2 |
|  | 1 kg | 0,50 mg | Ε2 |
|  | 2 kg | 1,0 mg | Ε2 |
|  | 5 kg | 2,5 mg | Ε2 |
|  | 10 kg | 10 mg | F1 |
|  | 20 kg | 30 mg | F1 |
|  | 50 kg | 500 mg | M1 |
| Mass / Non-automatic electronic weighing instruments | 1 mg to 5 mg | 0,0023 mg | Calibration according to guide EURAMET /cg18/ v.03 (2015). Available standard weights OIML R 111-1:2004: Class E1 at range 1 mg to 200 g Max grouped load available 411,11 g Class E2 at range 1 mg to 1000 g Max grouped load available 2111,11 g Class F1 at range 1 g to 5000 g Max grouped load available 11110 g Class F2 at range 10 kg to 20 kg Max grouped load available 50 kg Class M1 at range 200 g to 20 kg Max grouped load available 803,7 kg Calibration is performed on site |
|  | 10 mg | 0,0025 mg |  |
|  | 20 mg | 0,0027 mg |  |
|  | 50 mg | 0,0030 mg |  |
|  | 100 mg | 0,0035 mg |  |
|  | 200 mg | 0,0043 mg |  |
|  | 500 mg | 0,0053 mg |  |
|  | 1 g | 0,0063 mg |  |
|  | 2 g | 0,0080 mg |  |
|  | 5 g | 0,010 mg |  |
|  | 10 g | 0,024 mg |  |
|  | 20 g | 0,026 mg |  |
|  | 50 g | 0,028 mg |  |
|  | 100 g | 0,036 mg |  |
|  | 200 g | 0,063 mg |  |
|  | (200 g to 1 kg] | 1,0 · 10-6 |  |
|  | (1 kg to 10 kg] | 2,0 · 10-6 |  |
|  | (10 kg to 60 kg] | 1,5 · 10-5 |  |
|  | (60 kg to 860 kg] | 2,0 · 10-4 |  |
| Spectrophotometric Measurements |
| Absorbance / Absorbance of visible and UV radiation of spectrophotometers, in the range 200nm…900nm | to 0,5 AU | 6,7 mAU** | ASTM E 275: 2008 (re 2022) European Pharmacopοeia Edition 11, 2023 §2.2.25. (photometric accuracy for all wavelengths in the 200nm-880nm range & repeatability). Calibration can be performed on site. |
|  | (0,5 to 1,0] AU | 7,6 mAU |  |
|  | (1,0 to 1,5] AU | 8,4 mAU |  |
| Density Measurements |
| Digital density meters | 600 kg/m3 to 2000 kg/m3 | 0,006 % | According to SIM MWG7/cg-02/v.00, 2016 για mpe > 0.05 kg/m3 Calibration can be performed on site. |
| Digital alcoholmeters | 0 % to 100 % vol. EtOH | 0,012 % vol. EtOH |  |
| Density / Hydrometers, Areometers, Density Hydrometers, | 600 kg/m3 to 2000 kg/m3 | 0,28 kg/m3 | Calibration according the Cuckow method. |
| Specific Gravity Hydrometers | 0,6 to 2 | 0,28 ∙ 10-3 |  |
| Alcoholmeters | 0 % vol. EtOH to 100 % vol. EtOH | 0,30 % vol. EtOH |  |
| Light Baume Hydrometers | 10 ºBė to 103,3 ºBė | 0,08 ºBė |  |
| Heavy Baume Hydrometers | 0 ºBė to 72,5 ºBė | 0,06 ºBė |  |
| Density of liquids | 600 kg/m3 to 2000 kg/m3 | 0,16 kg/m3 | Internal procedure tQ-DENSITY-01 Calibration using hydrostatic weighing in thermostatically controlled set-up. |
| Density of liquids | 600 kg/m3 to 2000 kg/m3 +5 ºC to +80 ºC | 0,0030 % | Internal procedure tQ-DENSITY-02 based on ASTM D1480-15 and ASTM D1217-20 |
| Density of water – ethanol mixtures | 998,2 kg/m3 to 789,2 kg/m3 respectively 0 to 100% (v/v) EtOH, at 20 ºC | 0,010 % |  |
| pH measurements |
| pH meters, Ion meters | -410 mV to +410 mV | 0,035 mV | Internal procedure tQ-PH-01 Using a voltage simulator. Calibration can be performed on site. |
| Frequency / Time measurements |
| Revolution frequency / Centrifuges, Stirrers, Rotating Devices | [60 to 1.000] rpm | 1,2 rpm | Internal procedure tQ-TACHO-02 Using standard tachometer. Calibration can be performed on site. |
|  | (1.000 to 25.000] rpm | 2,0 rpm |  |
|  | (25.000 to 60.000] rpm | 5,0 rpm |  |
| Frequency / Optical Tachometers | [60 to 1.000) rpm | 0,5 rpm | Internal procedure tQ-TACHO-01 Using standard multifunction calibrator and standard tachometer. |
|  | [1.000 to 60.000] rpm | 1,0 rpm |  |
| Time / Stopwatches, Timers | (0 s to 1 h] | 0,06 s | Internal procedure tQ-TIME-01 Using standard stopwatch. Calibration can be performed on site. |
|  | (1 h to 5 h] | 0,24 s |  |
|  | (5 h to 10 h] | 0,48 s |  |
| Relative humidity measurements |
| Relative humidity / Analog and digital hygrometers, relative humidity data loggers | [25 to 50) %rh at 23 °C, 30 °C and 40 °C | 1,0 %rh | Internal procedure tQ-HYGRO-01 Comparative calibration using climatic chamber and standard hygrometer. |
|  | [50 to 80) %rh at 23 °C, 30 °C and 40 °C | 1,2 %rh |  |
|  | [80 to 90] %rh at 23 °C, 30 °C and 40 °C | 1,7 %rh |  |
| Relative humidity / Environmental chambers | [25 to 50) %rh at 23 °C, 30 °C and 40 °C | 1,2 %rh | According to guide DKD-R 5-7 09/2018, using standard relative humidity recorders. Calibration is carried out in the useful volume of the chamber or at selected points. Calibration can be performed on site. |
|  | [50 to 80) %rh at 23 °C, 30 °C and 40 °C | 1,4 %rh |  |
|  | (80 to 90] %rh at 23 °C, 30 °C and 40 °C | 1,9 %rh |  |
| Temperature Measurements |
| Temperature/ Climatic and isothermal chambers, ovens of column chromatography, ovens and incubators, refrigerators, freezers, ultra freezers, of capacity up to 2000 l, with or without air circulation | [-80 to -20) oC | 0,50 oC | According to guide DKD-R 5-7 09/2018, using Pt100 platinum resistance thermometer. Calibration is carried out in the useful volume of the chamber or at selected points. Calibration can be performed on site. |
|  | [-20 to +60) oC | 0,15 oC |  |
|  | [+60 to +250) oC | 0,20 oC |  |
| Temperature/ Climatic and isothermal chambers, ovens and incubators, refrigerators, freezers, ultra freezers, of capacity > 2000 l, with or without air circulation | [-40 to +85] oC | 0,60 oC | According to guide DKD-R 5-7 09/2018, using air temperature data loggers. Calibration is carried out on the useful volume or at selected points in the chamber Calibration can also be performed on site |
| Autoclaves | [+100 to +150] oC | 0,30 oC | According to guide DKD-R 5-7 09/2018, using Pt100 and Pt1000 platinum resistance thermometer. Calibration is carried out in the useful volume or at selected points in the chamber. Calibration can be performed on site. |
| Temperature/ Climatic and isothermal chambers, ovens of column chromatography, ovens and incubators, autoclaves, refrigerators, freezers of capacity up to 2000 l, with or without air circulation | [-20 to +150] oC | 0,50 oC | According to guide DKD-R 5-7 09/2018, using type K thermocouples. Calibration is carried out at selected points in the chamber, according to method (C) described in the technical guide DKD-R 5-7. Calibration can be performed on site. |
|  | (+150 to +350] oC | 0,80 oC |  |
| Temperature/ Climatic and isothermal chambers, ovens, of capacity up to 2000 L | (+350 to +500] oC | 1,3 oC | Internal procedure tQ-CHAMBERS-01 Method based on guide DKD-R 5-7 09/2018 (method C), using type K thermocouples. Calibration is carried out at selected points in the chamber. Calibration can be performed on site. |
|  | (+500 to +650] oC | 1,3 oC |  |
|  | (+650 to +1200] oC | 1,9 oC |  |
| Temperature/ Polymerase chain reactions (PCRs) | [0 to +30] oC | 0,50 oC | Internal procedure tQ-BLOCKS-01 Dynamic calibration throughout the plateau (16 measurements points) or at selected points, with the use of K type thermocouples. Heating/cooling and stabilized temperature uniformity, accuracy measurement. Overshoot/undershoot and ramp rate measurement. Calibration can be performed on site. |
|  | (+30 to +100] oC | 0,25 oC |  |
| Temperature/ Isothermal blocks, thermoreactors, COD testing instruments | [-40 to +150] oC | 0,10 oC | Εσωτερική μέθοδος tQ-BLOCKS-01 Με τη χρήση αισθητήρων αντίστασης πλατίνας Pt100. Η διακρίβωση διεξάγεται σε ολόκληρο το block ή σε επιλεγμένα σημεία. Η διακρίβωση μπορεί να γίνει και επί τόπου. |
|  | (+150 to +250] oC | 0,15 oC |  |
|  | [-40 to +150] oC | 0,20 oC | Internal procedure tQ-BLOCKS-01 Using platinum resistance thermometers, Pt100. Calibration is carried out in the whole block or at selected points of the block. Calibration can be performed on site. |
|  | (+150 to +250] oC | 0,30 oC |  |
| Temperature/ Isothermal plates | [-40 to 0) oC | 0,55 oC | Internal procedure tQ-PLATES-01 Using type K thermocouples. Calibration is carried out on the whole plate or at selected points of the plate. Calibration can be performed on site. |
|  | [0 to +40] oC | 0,37 oC |  |
|  | (+40 to +80] oC | 0,59 oC |  |
|  | (+80 to +150] oC | 1,8 oC |  |
| Temperature/ Baths up to 100 L | [-80 to +150] oC | 0,10 oC | Internal procedure tQ-BATH-01 Using Pt100 platinum resistance thermometer. Calibration is carried out in the useful volume or at selected points in the bath. Calibration can also be performed on site. |
|  | (+150 to +300] oC | 0,50 oC | Internal procedure tQ-BATH-01 Using type K thermocouples. Calibration is carried out at selected points in the bath. Calibration can also be performed on site. |
| Temperature/ Digital / analogue direct reading thermometers, digital temperature data loggers | [-80 to -20) oC | 0,030 oC | Internal procedure tQ-THERMO-01 Comparative method within isothermal bath using Pt100 platinum resistance thermometer |
|  | [-20 to +150] oC | 0,020 oC |  |
|  | [-35 to +80] oC | 0,17 oC | Internal procedure tQ-THERMO-03 Comparative method with dry block calibrator using Pt100 platinum resistance thermometer. Calibration can also be performed on site. |
|  | (+80 to +165] oC | 0,21 oC |  |
|  | (+165 to +270] oC | 0,10 oC |  |
|  | (+270 to +650] oC | 0,26 oC |  |
|  | (+650 to +900] oC | 1,2 oC | Internal procedure tQ-THERMO-03 Comparative method with dry block calibrator using type R thermocouples. Calibration can also be performed on site. |
|  | (+900 to +1200] oC | 1,4 oC |  |
|  | [-35 to 0) oC | 0,13 oC | Internal procedure tQ-THERMO-01 Comparative method within portable isothermal bath using Pt100 platinum resistance thermometer. Calibration can also be performed on site. |
|  | [0 to +80] oC | 0,06 oC |  |
|  | (+80 to +155] oC | 0,15 oC |  |
| Temperature/ Partial/total immersion, liquid in glass thermometers | [-80 to -20) oC | 0,03 oC | Internal procedure tQ-THERMO-01 Comparative method within isothermal bath with immersion depth 16cm using Pt100 platinum resistance thermometer. |
|  | [-20 to +150] oC | 0,03 oC | Internal procedure tQ-THERMO-01 Comparative method within isothermal bath with immersion depth 25cm using Pt100 platinum resistance thermometer. |
| Temperature/ Direct reading thermometers Digital/analogue temperature data loggers | [-40 to 0) oC | 0,20 oC | Internal procedure tQ-THERMO-02 Comparative method within temperature chamber using Pt100 platinum resistance thermometer |
|  | [0 to +45] oC | 0,10 oC |  |
|  | (+45 to +65] oC | 0,25 oC |  |
|  | (+65 to +100] oC | 0,50 oC |  |
|  | (+100 to +150] oC | 1,0 oC |  |
| Temperature/ Direct reading optical thermometers (Infrared thermometers) | [-10 to +50] oC | 0,30 oC | According to guide VDI/VDE 3511 part 4.4 : 2005 using black body source and Pt100 platinum resistance thermometer. |
|  | (+50 to +80] oC | 0,40 oC |  |
| Pressure Measurements |
| Relative Pressure / Relative pressure measuring equipment (gas medium) direct reading, analog – digital, pressure recorders | [-95 to +60] kPa | 15 Pa | Calibration according to guide DKD-R 6-1 03/2014. Calibration can be performed on site. |
|  | (+0,06 to +2,5] MPa | 0,45 kPa |  |
|  | (+2,5 to 10] MPa | 2 kPa |  |
| Absolute Pressure / Absolute pressure measuring equipment (gas medium) direct reading, analog – digital, pressure recorders | [+5 to +160] kPa | 15 Pa |  |
|  | (+0,16 to +2,6] MPa | 0,45 kPa |  |
|  | (+2,6 to 10,1] MPa | 2 kPa |  |
| Differential Pressure / Differential pressure measuring equipment (gas medium) | [-1 to +1] kPa | 2,0 Pa |  |
|  | [-20 to -1) kPa & (+1 to +20] kPa | 90 Pa |  |
| Barometric Pressure / Barometric pressure measuring equipment, direct reading, analog - digital, pressure recorders | [+800 to +1200] hPa | 0,4 hPa | Calibration according to guide DKD-R 6-1 03/2014. |
| Electrical Measurements |
| DC Voltage (Generation) / DC Voltage Measuring Equipment | [0 to 200] mV | 4,2 μV + 23·10-6·V | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | (0,2 to 2] V | 10,2 μV + 22,9·10-6·V |  |
|  | (2 to 20] V | 92 μV + 18,3·10-6·V |  |
|  | (20 to 200] V | 1,33 mV + 21,2·10-6·V |  |
|  | (200 to 1025] V | 11,6 mV + 21·10-6·V |  |
| AC Voltage (Generation) / AC Voltage Measuring Equipment | (0 to 200] mV |  | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | [10 to 45) Hz | 90 μV + 15,5·10-4·V |  |
|  | [45 to 1000) Hz | 74 μV + 1,8·10-4·V |  |
|  | [1 to 20) kHz | 81 μV + 4,97·10-4·V |  |
|  | [20 to 100) kHz | 2,6 mV |  |
|  | [100 to 500] kHz | 3 mV |  |
|  | (0,2 to 2] V |  |  |
|  | [10 to 45) Hz | 0,66 mV + 15,8·10-4·V |  |
|  | [45 to 1000) Hz | 0,68 mV + 1,4·10-4·V |  |
|  | [1 to 20) kHz | 0,67 mV + 3,8·10-4·V |  |
|  | [20 to 100) kHz | 26 mV |  |
|  | [100 to 500] kHz | 28 mV |  |
|  | (2 to 20] V |  |  |
|  | [10 to 45) Hz | 6,7 mV + 15,7·10-4·V |  |
|  | [45 to 1000) Hz | 7,65 mV + 1,18·10-4·V |  |
|  | [1 to 20) kHz | 7,63 mV + 2,7·10-4·V |  |
|  | [20 to 100] kHz | 45 mV + 15,1·10-4·V |  |
|  | (20 to 200] V |  |  |
|  | [10 to 45) Hz | 200 mV |  |
|  | [45 to 1000) Hz | 180 mV |  |
|  | [1 to 20] kHz | 174 mV + 3,3·10-4·V |  |
|  | (200 to 1020] V |  |  |
|  | [10 to 45) Hz | 518 mV + 1,6·10-4·V |  |
|  | [45 to 1000) Hz | 530 mV + 0,9·10-4·V |  |
|  | [1 to 10] kHz | 492 mV + 3,93·10-4·V |  |
| DC Current (Generation) / DC Current Measuring Equipment) | [0 to 200] μA | 35 nA + 75·10-6·I | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | (0,2 to 2] mA | 57,2 nA + 60,9·10-6·I |  |
|  | (2 to 20] mA | 0,44 μA + 34,1·10-6·I |  |
|  | (20 to 200] mA | 6,5 μA + 37,5·10-6·I |  |
|  | (0,2 to 2] A | 0,108 mA + 56·10-6·I |  |
|  | (2 to 30] A | 5,04 mA + 172·10-6·I |  |
| DC Current (Generation) / DC Current Clamp Meters | [1 to 60] A | 0,0082 A + 5,6·10-3·I | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | (60 to 220] A | 0,15 A + 6,7·10-3·I |  |
|  | (220 to 1500] A | 0,57 A + 5,3·10-3·I |  |
| AC Current (Generation) / AC Current Measuring Equipment) | [20 to 200] μA |  | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | [10 to 45) Hz | 393 nA + 1,34·10-3·I |  |
|  | [45 to 1000) Hz | 460 nA |  |
|  | [1 to 10] kHz | 0,36 μA + 10·10-3·I |  |
|  | (0,2 to 2] mA |  |  |
|  | [10 to 45) Hz | 2,7 μA + 0,95·10-3·I |  |
|  | [45 to 1000) Hz | 3 μA |  |
|  | [1 to 10] kHz | 2,3 μA + 4,87·10-3·I |  |
|  | (2 to 20] mA |  |  |
|  | [10 to 45) Hz | 25,9 μA + 1,01·10-3·I |  |
|  | [45 to 1000) Hz | 29 μA |  |
|  | [1 to 10] kHz | 24,6 μA + 3,08·10-3·I |  |
|  | (20 to 200] mA |  |  |
|  | [10 to 45) Hz | 260 μA + 1·10-3·I |  |
|  | [45 to 1000) Hz | 290 μA |  |
|  | [1 to 10] kHz | 250 μA + 3,05·10-3·I |  |
|  | (0,2 to 2] A |  |  |
|  | [10 to 45) Hz | 2,6 mA + 0,95·10-3·I |  |
|  | [45 to 1000) Hz | 3 mA |  |
|  | [1 to 5] kHz | 2,4 mA + 3,9·10-3·I |  |
|  | (2 to 30] A |  |  |
|  | [10 to 45) Hz | 26 mA + 0,97·10-3·I |  |
|  | [45 to 99) Hz | 23,1 mA + 1,87·10-3·I |  |
|  | [100 to 1000] Hz | 25,3 mA + 1,8·10-3·I |  |
| AC Current (Generation) / AC Current Clamp Meters (30 – 60Hz) | [1 to 60] A | 0,0082 A + 5,7·10-3·I | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | (60 to 220] A | 0,21 A + 6,9·10-3·I |  |
|  | (220 to 1500] A | 1,2 A + 5·10-3·I |  |
| 2-wire Resistance (Generation) / Resistance Measuring Equipment | 0,2 Ω | 5,9 mΩ | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | 0,3 Ω | 5,9 mΩ |  |
|  | 1 Ω | 5,9 mΩ |  |
|  | 10 Ω | 6,7 mΩ |  |
|  | 100 Ω | 10 mΩ |  |
|  | 1 kΩ | 0,079 Ω |  |
|  | 10 kΩ | 0,79 Ω |  |
|  | 100 kΩ | 7,4 Ω |  |
|  | 1 MΩ | 0,13 kΩ |  |
|  | 10 MΩ | 3,2 kΩ |  |
|  | 100 MΩ | 370 kΩ |  |
|  | 1 GΩ | 8,3 MΩ |  |
| 4-wire Resistance (Generation) / Resistance Measuring Equipment | 100 mΩ | 5,8 mΩ | Calibration according to guide EURAMET cg-15 v.3.0 : 2015. Using standard multifunction calibrator. |
|  | 1 Ω | 5,9 mΩ |  |
|  | 10 Ω | 6,6 mΩ |  |
|  | 100 Ω | 9,9 mΩ |  |
|  | 1 kΩ | 0,079 Ω |  |
|  | 10 kΩ | 0,79 Ω |  |
|  | 100 kΩ | 7,4 Ω |  |

*   	Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

** Uncertainty values were estimated assuming a spectrophotometer having a resolution of 0,1 mAU..

Site of assessment: Permanent laboratory premises, 20, Korgialeniou Str., Athens, Greece.

Approved Signatories: Konstantinos Chatzoudis, Georgios Chatzoudis. Dimitris Tentas, Eleni Chatzoudi, Chrisostomos Siope, Vasiliki Paraskevopoulou.

This Scope of Accreditation replaces the previous one dated 3.11.2020.

The Accreditation Certificate No. 426-6, to ELOT EN ISO/IEC 17025:2017, is valid until 06.05.2024.

Athens, February 9th, 2024

Christos Nestoras

CEO of ESYD

Ημερομηνία προτύπου. Διορθώθηκε.

Το πρότυπο δεν είναι πλήρες. Λείπει το Part και η ημερομηνία. Π.χ. VDI/VDE 3511 part 4.4 : 2005