Hellenic Accreditation System

Annex F2/13 to the certificate No. 500-5 SCOPE of ACCREDITATION of the Calibration Laboratory of “AUTOVISION SAKAR S.A. Private Technical Inspection of Vehicles - Certifications”

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2) * | Remarks |
| --- | --- | --- | --- |
| Dimensional measurements |
| Motor vehicles side slip measurement devices | (-15 to +15) m/km |  | Calibration according to manufacturer’s inspection instructions. Calibration is performed on-site. |
|  |  | 0,09 m/km | Actia Automotive S.A. (Actia Muller S.A., Muller BEM). Internal method ΓΔ.ΕΔ.020. |
|  |  | 0,12 m/km | Beissbarth GmbH. Internal method ΓΔ.ΕΔ.020B. |
|  |  | 0,20 m/km | MAHA Mashinenbau Haldenwang GmbH. Internal method ΓΔ.ΕΔ.020M. |
| Motor vehicles tires tread wear gauges | 0 mm to 3 mm | 0,17 mm | Internal method ΓΔ.ΕΔ.020. Calibration can also be performed on-site. |
| Motor vehicles headlights beam vertical inclination measurement devices | 0 % to 4 % | 0,12 % | Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020. Calibration can also be performed on-site. |
| Force measurements |
| Two-wheeled motor vehicles brake force measurement devices | [0 N έως 0,5 kN) [0,5 kN έως 1 kN) [1 kN έως 2 kN) [2 kN έως 2,7 kN] | 7 N 8 N 11 N 14 N | Manufacturer: Actia Automotive (Actia Muller, Muller). Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020. Calibration is performed on-site. |
| Light motor vehicles brake force measurement devices | [0 N έως 1,1 kN) [1,1 kN έως 2,2 kN) [2,2 kN έως 3,3 kN) [3,3 kN έως 4,4 kN) [4,4 kN έως 5,5 kN] | 6,6 N 7,4 N 9,4 N 12,0 N 15,6 N |  |
| Heavy motor vehicles brake force measurement devices | [0 N έως 11,1 kN) [11,1 kN έως 22,2 kN) [22,2 kN έως 33,3 kN) [33,3 kN έως 44,4 kN] | 59 N 0,11 kN 0,16 kN 0,21 kN |  |
| Light motor vehicles brake force measurement devices | [0 N έως 1,2 kN) [1,2 kN έως 2,5 kN) [2,5 kN έως 3,7 kN) [3,7 kN έως 5,0 kN) [5,0 kN έως 6,2 kN] | 25 N 22 N 24 N 25 N 23 N | Manufacturer: Beissbarth. Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020B. Calibration is performed on-site. |
| Heavy motor vehicles brake force measurement devices | [0 N έως 12,5 kN) [12,5 kN έως 31,0 kN) [31,0 kN έως 40,0 kN] | 0,31 kN 0,32 kN 0,33 kN |  |
| Two-wheeled motor vehicles brake force measurement devices | [0 Ν έως 1,5 kN) [1,5 kN έως 2,5 kN] | 12 N 13 N | Manufacturer: MAHA. Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020M. Calibration is performed on-site. |
| Light motor vehicles brake force measurement devices | [0 N έως 3 kN) [3 kN έως 4 kN) [4 kN έως 5 kN) [5 kN έως 6 kN] | 32 N 33 N 34 N 35 N |  |
| Heavy motor vehicles brake force measurement devices | [0 N έως 3 kN) [3,0 kN έως 3,75 kN) [3,75 kN έως 6 kN) [6,0 kN έως 10,0 kN) [10,0 kN έως 12,5 kN) [12,5 kN έως 20,0 kN) [20,0 kN έως 30,0 kN] | 74 N 77 N 87 N 0,11 kN 0,13 kN 0,19 kN 0,28 kN |  |
| Motor vehicles suspension testers |  | Calibration according to manufacturer’s inspection instructions. Calibration is performed on-site. |
|  | 0 N to 5900 N per wheel | 17 N | Actia Automotive S.A. (Actia Muller S.A., Muller BEM). Internal method ΓΔ.ΕΔ.020. |
|  |  | 45 N | Beissbarth GmbH. Internal method ΓΔ.ΕΔ.020B. |
|  |  | 32 N | MAHA Mashinenbau Haldenwang GmbH. Internal method ΓΔ.ΕΔ.020M. |
| Two-wheeled motor vehicles weighing devices | 1470 N to 4420 N | 12 N | Calibration according to manufacturer’s inspection instructions. Manufacturer: Actia Automotive (Actia Muller, Muller). Internal method ΓΔ.ΕΔ.020. MAHA Mashinenbau Haldenwang GmbH. Internal method ΓΔ.ΕΔ.020M. Calibration is performed on-site. |
| Photometric measurements |
| Motor vehicles headlight beam illuminance measurement devices | 6 lx to 144 lx (passing beams) 32 lx to 240 lx (driving beams) | 2,5 lx 7,7 lx | Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020. Calibration is performed on-site. |
| Exhaust emission measurements |
| Exhaust emission analyzers of class 0, I, and II, for motor vehicles operating with petrol | [CO] (0,5 to 5) % vol | 94 x 10-4 % vol | Calibration according to manufacturer’s inspection instructions. Internal method ΓΔ.ΕΔ.020. Maximum tolerance errors according to recommendation OIML R 99-1&2:2008, par. 5.5.2, table 4. Calibration can be performed on-site. |
|  | [CO2] (4 to 16) % vol | 62 x 10-3 % vol |  |
|  | [C3H8] (200 to 4000) parts per 106 vol | 11 parts per 106 vol |  |
| Exhaust emission analyzers for motor vehicles operating with diesel | Opacity (Ν) 0% to 100% Absorption coefficient (k) 0 m-1 to 9,99 m-1 | 0,69 % 0,016 m-1 | Calibration according to Directive 72/306/ΕOΚ, Annex VII, par 3.6. Internal method ΓΔ.ΕΔ.020. Calibration can be performed on-site. |
| Pressure measurements |
| Motor vehicles tires pressure measurement devices | 100 kPa to 1100 kPa | 14 kPa | According to Guide DKD-R 6-1:2014. Calibration can be performed on-site. |
| Pressure measurement devices of compressed air brake systems of heavy motor vehicles | 0 kPa to 1200 kPa | 10 kPa |  |

*	Where the expanded uncertainty (with coverage probability 95 %) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration & Measurement Capability (CMC) includes the measurand, the measurement range and the measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved in a calibration.

Permanent laboratory premises address: 77, Rovertou Galli str., 163 46 Ilioupoli, Greece.

Approved Signatory: D. Angelakopoulos

This Scope of Accreditation replaces the previous one dated 26.05.2023.

The Accreditation Certificate No. 500-5, to ELOT ΕΝ ISO/IEC 17025:2017, is valid until 19.01.2025.

Athens, 01.04.2024

Christos Nestoras

CEO of ESYD