Hellenic Accreditation System

Annex F2/7 to the Certificate No. 1120-3 SCOPE of ACCREDITATION of the Laboratory of Thermal Turbomachines (LTT) of the School of Mechanical Engineering of the National Technical University of Athens for the performance of calibrations

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Flow rate measurements |
| Gas flow meters of type: Diaphragm Rotary Turbine |  |  | Calibration according to: OIML R 137:2012 (Ε) EN 1359: 2017 EN 12480:2018 EN 12261:2018 Calibration medium: air at atmospheric conditions. |
|  | 0,006 m3/h to 0,04 m3/h | 0,63 % | Using drum type gas flow master meter. |
|  | 0,04 m3/h to 0,1 m3/h | 0,79 % |  |
|  | 0,1 m3/h to 0,36 m3/h | 0,38 % |  |
|  | 0,16 m3/h to 0,25 m3/h | 0,65 % | Using rotary type gas flow master meters. |
|  | 0,25 m3/h to 0,65 m3/h | 0,50 % |  |
|  | 0,65 m3/h to 1600 m3/h | 0,35 % |  |

* Where uncertainty (with a probability of 95% coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Calibration & Measurement Capability (CMC), includes the measured quantity, measurement range and measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved during a calibration.

Permanent laboratory premises address: 9, Iroon Politechniou str., 157 80 Zografou Campus, Zografou, Attiki, Greece.

Approved signatories: Konstantinos Mathioudakis, Nikolaos Aretakis.

This Scope of Accreditation replaces the previous one dated 11.10.2022.

The Accreditation Certificate No.1120-2 to ELOT EN ISO/IEC 17025:2017, is valid until 28.03.2026.

Athens, 25.01.2024

Christos Nestoras

CEO of ESYD