Hellenic Accreditation System

Annex F2/6 to the Certificate No. 1197-2

SCOPE of ACCREDITATION of the Calibration Laboratory of PROTEM GROUP SA

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2) * | Remarks |
| --- | --- | --- | --- |
|  |  | Static calibration according to Measurement and Control Institute guide WG00496: “A code of practice for the calibration of industrial process Weighing Systems”, 2011. Calibration is performed on-site. |
| Automatic weighing systems of industrial processes | (1 to 1250] kg | 2,7 x 10-4 | Using OIML Class M2 standard weights of range 0,1 kg to 20 kg. Maximum available grouped load: 1252 kg. |
|  | (100 to 10000] kg |  | Using substitute loads. Ι: load in kg d: resolution in kg n: number of substitute loads |
| Non-automatic weighing instruments |  |  | According to guide EURAMET cg-18, v. 4.0:2015. Calibration is performed on-site. |
|  | [0,1 to 5) kg | 4,0 x 10-5 | Using OIML Class M2 standard weights of range 0,1 kg to 20 kg. Maximum available grouped load: 1252 kg. |
|  | [5 to 1250] kg | 2,7 x 10-4 |  |
|  | (100 to 10000] kg |  | Using substitute loads. Ι: load in kg d: resolution in kg n: number of substitute loads |
| Standards weights |  | According to guide OIML R111-1:2004. Calibration capability of standard weights up to OIML class: |
|  | 0,1 kg | 0,9 mg | M1 |
|  | 0,2 kg | 1,3 mg |  |
|  | 0,5 kg | 2,0 mg |  |
|  | 1 kg | 5,9 mg |  |
|  | 2 kg | 7,5 mg |  |
|  | 5 kg | 90 mg | M2 |
|  | 10 kg | 0,10 g | M1 |
|  | 20 kg | 0,18 g |  |
| Compression testing machines | [200 to 3000] KN | 0,12 % | According to ISO 7500-1:2018. Calibration capability class 1 machines. Calibration is performed on-site. |
|  |  | Vertical force measurement: 0,12 % Force transfer: 1,0 % | According to ΕΝ 12390-4:2019. Calibration capability class 1 machines. Calibration is performed on-site. |

*	Where the expanded uncertainty (with coverage probability 95 %) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration & Measurement Capability (CMC) includes the measurand, the measurement range and the measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved in a calibration.

Permanent laboratory premises address: 35th road 3, 13341 Ano Liosia

Approved Signatories: G. Panagiotidis, D. Tselepis.

This scope of Accreditation replaces the previous one dated 06.02.2024.

The Accreditation Certificate No. 1197-2, to ELOT EN ISO/IEC 17025:2017, is valid until 18.02.2029.

Athens, 19.02.2024

Christos Nestoras

CEO of ESYD

# Text boxes

Mass measurements

Force measurements