Hellenic Accreditation System

Annex F2/1 to the Certificate No. 1237 SCOPE of ACCREDITATION of the Calibration Laboratory Of Emmanouil C. Manolas (EL.E.PI.)

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Measurements |
| Mass / Standard weights |  | Calibration of standard weights according to OIML R 111-1:2004 (Ε). Calibration capability up to (and including) OIML Class : |
|  | 1 kg | 160 mg | M3 |
|  | 2 kg | 160 mg | M3 |
|  | 5 kg | 160 mg | M2 |
|  | 10 kg | 160 mg | M1 |
|  | 20 kg | 300 mg | M1 |
|  | 500 kg | 25 g | M2 |
| Mass / Non-automatic weighing instruments | 500 g | 27 mg | Calibration according to guide EURAMET /cg-18/v.04 (11/2015). Use of standard weights of OIML R 111 class: M1 in rage 500 g … 20 kg. Max grouped load: 650 kg M2 in rage 500 kg …19.5 t. Max grouped load: 19,5 t. Calibration is performed on-site. |
|  | 1 kg | 55 mg |  |
|  | 2 kg | 110 mg |  |
|  | 5 kg | 270 mg |  |
|  | 10 kg | 550 mg |  |
|  | 20 kg | 1,1 g |  |
|  | 500 kg | 100 g |  |
|  | 500 kg … 19500 kg | 2,1 x 10-4 |  |
|  | 0,6 t … 3 t | kg | Calibration according to guide EURAMET /cg-18/v.04 (11/2015), with use of substitution loads. Where: Ι, max load indication in kg, d, instrument's readability in kg and Ν, steps (number of substation loads + 1) Calibration is performed on-site. |
|  | 3 t … 60 t | kg |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Site of assessment: Permanent Laboratory premises, 2nd km N.R. Katerini - Thessaloniki, 60100, Katerini.

Approved Signatories: E. Manolas, P. Nasios.

The Accreditation Certificate No 1237, to ΕLΟΤ ISO/IEC 17025:2017, is valid until 9.02.2025.

Athens, 10th February 2021

# Spyridon Podaras

Managing Director of ESYD