Hellenic Accreditation System

Annex F2/9 to the Certificate No. 547-5 SCOPE of ACCREDITATION of the Renewable Energy Laboratory of INTERNATIONAL WINDENGINEERING G.P. for the performance of calibrations

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2) * | Remarks |
| --- | --- | --- | --- |
| Air speed measurements |
| Ταχύτητα αέρα / Cup anemometers, uniaxial propeller anemometers, ultrasonic anemometer, hot wires/films, air flow meters Directly indicating transmitters. Transmitters with analog output 0 … 20 mA, 0 … 10 V DC Transmitters with digital output | 4 m/s … 16 m/s | 0,13 m/s (at 4m/s) 0,08 m/s (at 10m/s) 0,10 m/s (at 16m/s) | IEC61400-12-1: 2017, Annex F Calibration is performed in a closed circuit wind tunnel |
|  | 1 m/s … 30 m/s | 0,19 m/s (at 1 m/s) 0.16 m/s (at 2m/s) (for speeds 4 – 16 m/s, see uncertainties above) 0.14 m/s (at 20m/s) 0.19 m/s (at 30m/s) | Calibration according to internal method CLP-07 based on: IEC61400-12-1: 2017, Annex F Calibration is performed in a closed circuit wind tunnel |
| Wind direction measurements |
| Wind direction / Wind direction sensors: Potentiometric sensors with voltage measurement Sensors with serial output Sensors connected to conditioning modules Sensors connected to digital/analogue devices (ultrasonic anemometers, hot-wire anemometers, propeller anemometers) | 0°… 360° | 0,8° | IEC61400-12-1:2017, Annex N Calibration is performed in a closed circuit wind tunnel |
| Temperature measurements |
| Temperature / Temperature transducers with analog output in mV or mA | -20oC…90oC | 38 mK | Comparative calibration according to internal method CLP-05 based on: DKD-R-5-1: 2018 using bath and reference Pt thermometer |
| Pressure measurements |
| Pressure / Pressure transducers with analog output in mV or mA | 500 … 1100 hPa | ±0,21 hPa | Comparative calibration according to internal method CLP-08 based on: DKD-R 6-1: 2014 using reference barometer and controlled pressure chamber |
| Electrical Measurements |
| DC Voltage / Recorders, multimeters, DC voltage measuring equipment | 1μV … 100mV | 0,0050%Voltage + 3,5 μV | Comparative calibration according to internal method CLP10-02 based on: EURAMET cg15, v.3: 2015 |
|  | >100mV … 1V | 0,0035%Voltage + 7 μV |  |
|  | >1V … 10V | 0,0030%Voltage + 0,05 mV |  |
|  | >10V … 30V | 0,0040%Voltage + 0,6 mV |  |
| DC Current / Recorders, multimeters, DC current measuring equipment | 100 μΑ | 0,05%Amperage + 0,025 μΑ |  |
|  | >100μΑ … 1 mA | 0,05%Amperage + 0,06 μΑ |  |
|  | >1mA … 10 mA | 0,05%Amperage + 2 μΑ |  |
|  | >10mA …20 mA | 0,05%Amperage + 5 μΑ |  |
| Frequency / Recorders, multimeters, DC frequency measuring equipment | 1 … 500 Ηz | 0,2 Hz |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

.

Site of assessment:  Permanent laboratory premises, Theotokopoulou 24, 153 44 Gerakas, Attiki, Greece and wind tunnel of NTUA

Approved signatories: E. Morfiadakis, K. Papadopoulos

This Scope of Accreditation replaces the previous one dated May 18th, 2021.

The Accreditation Certificate No. 547-5, to ELOT EN ISO/IEC 17025: 2017, is valid until 18.5.2025.

Athens,  July 28th, 2021

Spyridon Podaras

CEO of ESYD