Hellenic Accreditation System

Annex F2/14 to the Certificate No. 490-5 SCOPE of ACCREDITATION of the High Voltage Service Providing Laboratory Unit (SPLU) of the School of Electrical and Computer Engineering of the National Technical University of Athens for the performance of calibrations

| Measurand / Calibration item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Electrical Measurements |
| Electrostatic Discharge Generator (ESD) DC voltage Peak current Current at 30 ns Current at 60 ns Rise time | 100 V to 32 kV 0,5 A to 50 Α 0,5 A to 50 Α 0,5 A to 50 Α 500 ps to 10 ns | 1% 5% 5% 5% 9% | ELOT ΕΝ 61000-4-2 Ε2:2009 |
| Burst generator at 50Ω and 1 kΩ Peak voltage Rise time Impulse duration Frequency | 20 V to 6000 V 1 ns to 100 ns 10 ns to 10 s 1 Hz to 1 MHz | 5% 4% 2% 1% | ELOT ΕΝ 61000-4-4: E3: 2012 ISO 7637-2 E3:2011 ISO 7637-3 E3:2016 §4.5, §4.6 |
| Surge generator Peak voltage (1,2/50 μs) Front time Time to half value Impulse duration Peak current (8/20 μs) | 100 V to 7000 V 100 ns to 10 μs 1 μs to 100 μs 1 μs to 10 s 1 Α to 10000 Α | 2% 2% 2% 2% 2% | ELOT ΕΝ 61000-4-5 E3:2014 + A1:2017 ISO 7637-2 E3:2011 ISO 7637-3 E3:2016 §4.5, §4.6 ISO 16750-2:2012 §4.6.3 |
| Measurand / Calibration item | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| Voltage dips, short interruptions and voltage variations equipment Voltage Rise time Angle | 10 V to 400 V 0,1 ns to 100 ns 0° to 360° | 2,5% 10% 3,5% | ELOT ΕΝ 61000-4-11: 2004 + A1:2017 |
| Power frequency magnetic field immunity test equipment Current Magnetic flux density | 0.1 A to 16 A 0,1 μΤ to 500 μT | 2% 5% | ELOT ΕΝ 61000-4-8: 2010 |
| (Large Loop Antennas System – LLAS) Current | -120dBm to +20dbm | 2,2dB | ELOT EN 55016-1-4 Ε4:2019 + A1:2020 |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Site of assessment: Permanent Laboratory Premises, 9, Iroon Politechniou str., 157 80 Zografou Campus, Zografou, Attiki, Greece

Approved signatories: Ioannis F. Gonos

This Scope of Accreditation replaces the previous one dated December 21st,2022 .

The Accreditation Certificate No. 490-5, to ELOT EN ISO/IEC 17025: 2017, is valid until 14.12.2025.

Athens, March 15th, 2024

Christos Nestoras

CEO of ESYD