Hellenic Accreditation System

Annex F2/14 to the Certificate No. 705-5

SCOPE of ACCREDITATION of the Calibration Laboratory of the Val ELECTRONIC L.P.

| Calibration item/ Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Measurements |
| Non-automatic balances/ Mass | 1 mg | 0,004 mg | Calibration according to Technical Guide EURAMET cg-18 v.4.0:2015 Available OIML class standards: E1, range 1 mg to 50 g (max available load 100 g) E2, range 1 mg to 200 g (max available load 500 g) F1, range 1 g to 5 kg (max available load 10 kg) M1, range 5 kg to 200 kg All calibrations are performed on-site. |
|  | 2 mg | 0,004 mg |  |
|  | 5 mg | 0,004 mg |  |
|  | 10 mg | 0,004 mg |  |
|  | 20 mg | 0,004 mg |  |
|  | 50 mg | 0,005 mg |  |
|  | 100 mg | 0,006 mg |  |
|  | 200 mg | 0,007 mg |  |
|  | 500 mg | 0,010 mg |  |
|  | 1 g | 0,012 mg |  |
|  | 2 g | 0,015 mg |  |
|  | 5 g | 0,019 mg |  |
|  | 10 g | 0,024 mg |  |
|  | 20 g | 0,030 mg |  |
|  | 50 g | 0,037 mg |  |
|  | 100 g | 0,061 mg |  |
|  | 200 g | 0,37 mg |  |
|  | 500 g | 0,97 mg |  |
|  | 1 kg to 10 kg | 6,1 x 10-6 |  |
|  | 20 kg to 200 kg | 6,1 x 10-5 |  |
| Temperature Measurements |
| Direct reading thermometers/ Temperature: a) analog & digital b) total immersion, liquid in glass thermometers | [-25 oC to 150 oC] | 0,15 oC | Internal laboratory method OE12-1 «Thermometer Calibration Method» based on recommendation DKD R-5-1:2018 |
|  | 150 °C to 200 °C | 0,21 °C |  |
| Dimensional Measurements |
| Callipers/ Length: |  | VDI/VDE/DGQ 2618, part 9.1: 2006 |
| a) Analog / division 20 μm | Up to 200 mm | 11 μm |  |
| b) Analog / division 50 μm |  | 29 μm |  |
| c) Analog / division 100 μm |  | 57 μm |  |
| d) Digital / division 10 μm |  | 6,0 μm |  |
| Length / External spindle micrometers: |  | VDI/VDE/DGQ 2618, part 10.1: 2001 |
| a) Analog & Digital / division 10 μm | Up to 25 mm | 5,7 μm |  |
| b) Analog & Digital / division 1 μm |  | 0,60 μm |  |

* Where the expanded uncertainty (with 95 % coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration Measurement Capability (CMC) includes the measured quantity, the measurement range and the measurement uncertainty, expressing the minimum measurement uncertainty which can be achieved in a calibration.

Permanent laboratory premises address: Parthenonos 1-3, GR-146 71 Nea Erithrea, Attiki, Greece.

Approved Signatory: A. Valelis.

This Scope of Accreditation replaces the previous one dated 12.07.2023.

The Accreditation Certificate No. 705-5, to ELOT EN ISO/IEC 17025:2017, is valid until 13.09.2027.

Athens, 13.09.2023

Christos Nestoras

CEO of ESYD